## Blank
Test Requirement
Standard or ETP | IEC 62301 section 4 as referenced by both 10CFR430.23 subpartB App J1
Key Specs | (Identify specs defined by test standard or procedure)
Watt-hr to 0.1 resolution | 5.3.2 Sampling Method | 1.
Total Harmonic Content | Section 4.3.2 Supply Voltage Waveform | 2.
Crest Factor | Section 4.4.1 Power Measurement Uncertainty | 3.
Power Factor | Section 4.4.1 Power Measurement Uncertainty | 4.
Equipment evaluated
Yokogawa WT210 | We had 3 available WT210's with harmonics in AP3. These met the standard, therefore I used before buying more
Camille Bauer APLUS | The APLUS meets the minimum requirements, however does not meet the recommeded requirments (Crest Factor specifically) | 1395
Yokogawa WT332-D-C1/C7/G5 | The WT210 has been discontinued. The replacement is the WT332. I purchased 5 pieces. This device offers an additional advantage (not required by IEC 62301) of simultaneous Normal and Harmonic measurements | 5640
Equipment Selected
Yokogawa WT210 | Meets standard - Uses integrated watt hours
Yokogawa WT332-D-C1/C7/G5 | Exceeds standard - Uses integrated watt hours
Meets test standard requirement? (yes/no) |  |  | Yes
In scope of current calibration supplier(yes/no) |  |  | No
If No, what is calibration plan:
Calibrated by OEM. Calibration supplier is doing a scope change
Instrumentation Engineer/Lab Manager |  |  | Date
Corey Sweatt

## Example
Test Requirement | Precision Floor scale
Standard or ETP | DOE CFR, AHAM Wash Perf, Vent Condensate
Key Specs | (Identify specs defined by test standard or procedure)
Range up to 250 lbs |  | 1. | 500 lbs
Resolution 0.005 lbs (dry), 0.05 (wash) |  | 2. | .05 lbs
Accuracy+/- 0.05 lbs (dry), 0.1 (wash) |  | 3. | .1 lbs
24" x 24" base |  | 4. | 24"x24"
Equipment evaluated
American scale, Grainger, Premier Scale
Equipment Selected
Will trial , B-Tek Bench Scale, Model BT-BC-2424-50-SP
Meets test standard requirement? (yes/no) |  |  | Yes for wash
In scope of current calibration supplier(yes/no) |  |  | No
If No, what is calibration plan: |  | Premier Scale will calibrate
Instrumentation Engineer/Lab Manager |  |  | Date